## Overview
BTS-600 Program Editor — Feature List
Source: BTS-600 / Battery Manager User Manual (BM_Manual_eng.pdf) — Chapters 11.1, 11.2, 12 (BTS-600 Programs)
| Purpose | Reference feature inventory for the Program Editor module, to be used as the basis for the web app's functional spec / backlog.
| Scope | Program (test-sequence) creation, editing, and management for BTS-600 battery test systems: grid-based step editor, operators, nominal values, limits, actions, registration formats, special operators, cycles, procedures, and program-level management (create/save/delete/rename/print/lock).
| Sheet Guide | Tab 1: Program Management ¦ Tab 2: Editor Grid & Editing Functions ¦ Tab 3: Operators ¦ Tab 4: Nominal Values ¦ Tab 5: Limits ¦ Tab 6: Actions ¦ Tab 7: Registration & Formats ¦ Tab 8: Special Operators (Deep Dive) ¦ Tab 9: Cycles, Procedures & Counters ¦ Tab 10: Battery Parameters
| Total Sheets | 10 detail sheets, ~140 individual feature/operator/option rows

## 1. Program Management
Program Management
Navigation, create/save/delete/rename/print/lock of programs
Feature | Description | Sub-options / Values
Program listing / tree view | Left-panel tree of programs organized by folders/layers. Filter by name or number (matches highlighted with yellow background). "Preview" radio button switches to a sortable table of all programs and their details; column layout is auto-saved. | Sortable columns; filter by name/number
Folder management | Right-click a location to create a folder or subfolder. Programs can be dragged and dropped into folders for organization. | Up to 20 levels of nested subfolders
Open Program Editor | Double-click a program name in the tree, or select it and click "Edit Program" from the Program Information window, to open the step-grid editor.
Calculation / graph view | Button opens a graphical plot of the program's function (current on Y-axis, time on X-axis) for supported systems. Right-click the chart → Properties to customize its appearance.
Create new program | Right-click the battery-type node in the tree → "New Program". Enter a unique program name and select a free program number (manually or via "Next available Program Number"). Save via OK. | Program name must be unique per battery type — ambiguous naming can cause fatal test errors
Save program | Save via the Save button(s); "Save Program As" requires a name and a free program number. | "Next available Program Number" auto-assign option
Cancel / unsaved-changes prompt | Cancel terminates the editor; if there are unsaved changes, the user is prompted to save before closing.
Delete program | Right-click a program name in the tree → Delete.
Rename program | Right-click a program name → Rename; dialog allows changing both the program name and its comment.
Print program | File menu provides Print (direct to system print dialog), Print Setup (printer properties), and Print Preview (page layout preview with Next/Prev Page, Two-Page view, Zoom In/Out, Close). | Print / Print Setup / Print Preview / Zoom / Two-page view
Program locking ("Checked") | Restricts editing on a per-program basis. Requires enabling the "Certificator" role in password maintenance and setting Certificator=TRUE in BM.ini under [SYSTEM_EDITOR]. Adds a "Checked" checkbox to Program Information; once checked, the program becomes view-only. | Role-based permission (Certificator); config flag in BM.ini
Multi-system program support | The editor adapts its available fields and workflow depending on the target system type (e.g. test, formation, or laboratory system). | Test / Formation / Laboratory system variants

## 2. Editor Grid & Editing
Editor Grid & Editing Functions
Step-grid columns and in-editor editing operations
Feature | Description | Sub-options / Values
Editor Grid Columns
Step (column) | Auto-generated, sequential program step number.
Label (column) | Usually blank. Holds GOTO jump-target labels, comment markers (!), or special labels (ONERROR / ONEXIT). | GOTO label · ! comment/disable marker · ONERROR · ONEXIT
Operator (column) | Program command defining the step type (charge, discharge, pause, etc.) or the name of a reusable procedure. | See "3. Operators" sheet
Nominal Value (column) | Numeric value plus unit controlling the operator (e.g. charge current, voltage, power). | See "4. Nominal Values" sheet
Limit (column) | Condition that ends the step — time-based or channel-threshold based. | See "5. Limits" sheet
Action (column) | What happens when the Limit is reached (continue, interrupt, stop, jump, load procedure). | See "6. Actions" sheet
Registration (column) | Frequency/condition for recording measurement data for that step. | See "7. Registration & Formats" sheet
Channel-unit constraint | Channel units (e.g. A, V, C) can be used directly in Nominal Value / Limit / Registration columns, but channel names cannot be referenced directly in programs.
Editing Functions
Copy program steps | Highlight one or more contiguous steps by dragging in the Step column, then right-click → Copy. | Contiguous selection only
Paste steps | Right-click the destination step → Paste. Copied steps are inserted above the highlighted step, in the same program or another program of the same system.
Insert step(s) | Right-click a step and choose "Insert step(s)", or press the Insert key, to add new step(s) ahead of the current selection. A dialog specifies how many steps to add.
Edit field (input assist) | Clicking a field reveals an assist button that opens the appropriate helper: operator list, value-list dialog, or Pascal Editor depending on the column.
Delete field content | Backspace/Del clears the currently open field's content. Spacebar/Del on an unopened field opens and clears it in one action. Esc restores the field's previous value; Undo steps back through changes.
Delete step(s) | Highlight one or more steps, right-click → "Delete step(s)".
Comment lines (BTS-600 only) | Enter "!" in the Label column to insert a comment line (confirmed with arrow keys, not Enter); enter the comment text via the field's edit button. Multiple comment lines require separate rows. | ! marker
Deactivate / disable step | Enter "!" in the Label column of an existing step to temporarily disable it; the step is highlighted yellow. Removing "!" restores it to a normal, active step. | ! marker; visual: yellow highlight
Help Dialogs | Clicking the assist button in the Operator, Nominal Value, or Registration columns opens an alphabetical list of all valid selections for that column; selecting an entry (click + Enter/OK) replaces the field content.
Pascal Editor | Opened from Nominal Value or Limit columns to view/author Pascal-based nominal value functions. Replace/Add apply a function to the field; Show displays read-only Pascal source; Edit/Delete manage user functions; New creates a function. | S_USER functions are editable/deletable; S_DIGA functions are read-only and cannot be deleted, especially if already used in a saved program

## 3. Operators
Operators
Program commands entered in the Operator column
Feature | Description | Sub-options / Values
SET | Sets registration formats, user-defined variables/counters, RLevel, timers, and global limits.
PAU (Pause) | Drops power contactors; step runs until its Limit is reached.
CHA (Charge) | Charges with the given nominal value until the Limit is reached; sets the charge relay.
RCH (Recharge) | Same behavior as CHA, but also sets the recharge relay; charge LED blinks.
DCH (Discharge) | Discharges with the given nominal value until the Limit is reached; sets the discharge relay.
INT (Interrupt) | Interrupts the running program; resumed via the front-panel Start button or a host command.
STO (Stop) | Terminates the program or sub-program; can appear multiple times in a program.
BEG (Cycle Begin) | Marks the start of a repeatable cycle block; an optional cycle name can be given in Nominal Value.
CYC (Cycle End) | Marks the end of a cycle block; Nominal Value holds the number of repetitions.
GOTO | Jumps to a step whose Label matches the given destination (used as Operator with target in Nominal Value, or as an Action).
REG | Registers a single reading using the current (or a named) registration format.
TASK | Activates a parallel background process (e.g. relay or temperature control); up to 12 parallel processes supported. | OUW, CGRE, CLESS, TCONTR built-in tasks
RET (Return) | Returns from a procedure back to the calling program/level.
SETMUX (Switch) | Activates an additional multiplexer channel display for serial data acquisition.
FILE | Sets the file / test-section name to be used for subsequent registrations.
ADD | Adds together channel values entered in the Nominal Value column.
BATT | Battery Simulator preset: constant voltage plus upper/lower current limits. | Units: V, A, AG
ONERROR / ONEXIT | Special Label operators used as jump targets — ONERROR on any program error, ONEXIT after normal program completion. | See "9. Cycles, Procedures & Counters" sheet
PROT | Starts an external protocol/report program, passing circuit, battery, and session-ID as arguments. | Path configured in [MAIN] ReportExeFile of the BTS-600 custom INI
SUB | Subtracts the second channel value from the first; respects sign for A/mA limits when used.
PAUA, PAUO, URANGE, IRANGE, OUTA, OUTB, ISOEXT, ISOINT | Additional operators referenced but detailed only in the separate "BM_PM_BTS600_New_Operators" document (not covered in depth in this manual). | Follow up with the New Operators reference doc for full spec
RANGE / IRANGE | Controls current-range switching: automatic (default) or manual selection. | 1–4 manual ranges (1–2 for IGBT); 0 = automatic
PARALLEL | Links two or more circuits together for synchronized parallel test operation. | Up to 7 slave circuits on MBT master; up to 15 on ME master
EIS | Runs an EIS-Meter electrochemical impedance spectroscopy step. | See EIS parameters in "8. Special Operators" sheet
SYNCLine / SYNCProgram | Synchronizes circuits sharing the same SyncGroup at a given program step. | SYNCProgram requires identical program + version; SYNCLine ignores program identity
SAVE | Saves the current registration-format battery counter data into a named test section. | Supports multiple named sections per battery
REST | Restores counter data (Ah, Wh, etc.) previously stored by SAVE, to continue a test from saved state. | Section name must match; requires matching registration format pre-declared
TABLE | Loads and executes a nominal-value table/profile (drive-cycle style) from a text file. | Built-in profiles: FUDS, SFUDS, DST
CLEAR | Deletes SET registrations/limits and global setpoints; does not clear or deactivate the registration format itself.
FILTER | Sets digital filters per channel (Digatron-configured hardware only).

## 4. Nominal Values
Nominal Values / Functions
Values entered in the Nominal Value column
Feature | Description | Sub-options / Values
Standard functions | Basic control values for CHA/RCH/DCH steps: Current, Voltage, Power, Resistance. A step must always define current, power, or resistance — never voltage alone. | Units: A, V, Watt, Ohm (power must be entered as "Watt", not "W")
Battery-parameter-based functions | Derived nominal values based on battery parameters rather than fixed numbers, e.g. a C5 discharge-rate current. | ACn5, ACn1, ACn2, ACn4, ACn10, ACn20, VnC — derived from Nom. Capacity / Cells; not usable in ramps
Ramp functions | Define a start value, end value, and duration to linearly ramp a control value over time. | Applicable to current, voltage, power, resistance, or combinations; e.g. "1000-5000 Watt 300 sec"
Parallel function | Combine two nominal values in a single step (e.g. constant-voltage + constant-current) — neither value is exceeded.
Special nominal values | Purpose-built values for advanced control: Factor_I/Factor_P/Factor_U scale TABLE values by a factor; OUW activates resistance/power calculation (paired with TASK); CGRE/CLESS activate/deactivate a temperature-dependent relay via IOOUT[16]; TCONTR switches a relay based on temperature. | Factor_I, Factor_P, Factor_U, OUW, CGRE, CLESS, TCONTR
Reserved nominal values | System-reserved names used for specific mechanisms: File Name (TABLE/REG/FILE), Filter Name, TIMER1–TIMER3 (SET-started timers), MaxCHAI/MaxDCHI (current limits), MaxChaU/MaxDchU (voltage limits), MaxChaW/MaxDchW (power limits), Number (message ID for MSG/ERR). | File Name, Filter Name, TIMER1-3, MaxCHAI, MaxDCHI, MaxChaU, MaxDchU, MaxChaW, MaxDchW, Number

## 5. Limits
Limits
Termination conditions entered in the Limit column
Feature | Description | Sub-options / Values
Time limit | Ends the step after a fixed duration (step duration, not a clock time). | Formats: sec, min, h, or hh:mm:ss
Channel threshold limit | Ends the step when a physical or logical channel crosses a threshold. | < value UNIT (lower) / > value UNIT (upper); channels: A, V, Ah, Wh, WhDch, K, etc.
Linked (AND) limits | Chains two or more limit conditions across successive rows; the action fires only once all linked conditions are met, and must be attached to the last row in the chain. | '&' linkage operator
Timer1–Timer3 limit | Ends the step when a SET-started timer elapses. | Timer1, Timer2, Timer3
GradT / GradTm | Ends the step when temperature rise exceeds a threshold, measured per hour (GradT) or per minute (GradTm).
GradU(m) / GradI(m) | Ends the step when voltage/current rise exceeds a threshold per hour (or per minute with the 'm' suffix). Requires device adjustment; limited to 1 channel; negative sign detects a decrease instead of increase.
deltaV | Applies a dChannel/dt gradient algorithm to voltage or GREAL channels. Standard usage monitors a sense-input; enhanced usage adds gradient time-base analysis (ms) across multiple channels (GREAL[100]-[119]) via ChanIdx/ChanNum/TimeIdx globals.
AhDef | Marks the current Ah value as the 100% reference point; usable only with the PAU operator.
PercAh | Ends the step once Ah reaches X% of the AhDef reference value.
PERCCN_P | Ends the step when the previous step's ending capacity reaches X% of nominal capacity (CNom); sign is ignored.
PERCCN_C | Ends the step when the current step's capacity reaches X% of nominal capacity (CNom); sign is ignored.
A_notAbs / mA_notAbs | Sign-aware current limits that trigger only on positive current, preventing false trips from negative (discharge) current.
Special customer limit options | Derived/customer-specific channel-value limits. | ABATT, ACN5, OHM, VBATT, VNC, WATT

## 6. Actions
Actions
Behavior triggered when a step's Limit is reached
Feature | Description | Sub-options / Values
(blank) | No action entered — program simply continues to the next test step.
INT | Interrupts the program when the limit is reached; resumed via the Start button or a host command.
STO | Stops only the currently running program; does NOT stop the remainder of a dispo-list (queued) tests — flagged as a risk of unintended continued discharge.
GOTO (jump destination) | Jumps to the step whose Label matches the destination. Restricted to the same program level — cannot jump into or out of a procedure.
Procedure name (as action) | Loads and executes the named procedure, then continues at the next step.
ERR (as action) | Writes a numbered message to the registration data and interrupts the step; resumable at the same position. | See ERR message codes in "8. Special Operators" sheet
MSG (as action) | Writes a numbered message to the registration data without interrupting the program.

## 7. Registration & Formats
Registration & Registration Formats
Data-logging configuration entered in the Registration column
Feature | Description | Sub-options / Values
Registration format concept | Defines which measurement channels are recorded and in what order. Set via the SET operator (Operator column) with the format name entered in the Registration column.
Preset registration formats | Built-in named formats bundling a fixed set of channel units, each with a documented memory cost per registration. | CHANREG/CHAREG (42 byte), CYCLE/CYCLEC/DCHREG family (90 byte, incl. A,V,C,Ah,AhBal,AhStat,AhCha,AhDch,AhStep,AhPrev,Wh,WhCha,WhDch,WhStep,WhPrev,V1), GSM (A,A_Diff,A_Low,Ah,V,V_Drop,V_High,V_Low), SIMPLE (A,V — 12 byte), STANDARD (A,V,C,Ah,AhStep,Wh,WhStep — 42 byte)
Logical channel counters | Derived counters available for registration beyond raw physical channels. | Ah/Wh (total), AhStat/AhBal (charge-factor-adjusted), AhStep/WhStep (per-step), AhPrev/WhPrev (previous step), AhCha/WhCha (cumulative charge), AhDch/WhDch (cumulative discharge)
Create new registration format | "Create new registration format" in the help dialog; name the format, then move channel units from "All units" into the selected list.
Modify registration format | Add or remove channel units from an existing format; OK saves, Cancel discards. Editing a shared format affects every program that uses it.
Change format mid-program | A new SET operator can redefine the active format anywhere in the program (including inside a procedure, scoped to that sub-program level), or a single step's format can be overridden by entering a format name directly in that step's Registration column.
RLevel variable | Dynamically activates/deactivates registration recording via SET RLevel = value. | 0.0 all off · 1.0 step-changes only · 2.0 all (data+step-change) · 3.0 suppress step-change (data only at step end) · 4.0 suppress step-change (no end-of-step data either)
EIS registration format | Special-purpose format for EIS-Meter channels, defined via SET + Registration column after integrating EIS-Meter channels.
Registration frequency/trigger types | Controls how often/under what condition a registration entry is written for a step; multiple entries can be combined per step. | Delta (value-change threshold, e.g. "2.0 A") · Time interval (e.g. "5 min") · Threshold < / > (e.g. "<0.5 A", ">50.0 C") · Range >> / << with & (delayed-start delta, e.g. ">> 1h & 2.0V") · Maximum count * (fixed number of registrations, e.g. "15*", ">> 10000 Watt 5*")

## 8. Special Operators
Special Operators — Deep Dive
Detailed behavior of advanced operators (section 12.4.8)
Feature | Description | Sub-options / Values
SET (variables) | Assigns a user-defined variable; if the variable name matches a channel unit, that channel is set directly. Commonly used to preset cycle counters dynamically.
Global Nominal Values / Limits | A program-wide limit that no step may exceed, checked throughout the whole run. The paired action is restricted to ERR or MSG; if no action is set and the limit is reached, the entire program terminates.
REG (as program step) | Registers a single reading using the current (or an optionally named, in Nominal Value) registration format — useful for extracting e.g. per-cycle capacity into a separate section.
TASK | Activates a parallel background process; up to 12 concurrent tasks. Selected via a numbered help dialog. Must be paired after PARALLEL if using Safetask. | Built-in tasks: OUW (resistance/power calc), CGRE/CLESS (temp-relay on/off), TCONTR (temperature-based relay switching)
ERR | Interrupts the program and logs a numbered message; resumable at the same position. Default message codes are documented; message text is editable via Maintenance → BTS-600 Messages. | 1 Limit reached · 2 Voltage limit · 3 Temperature limit · 4 Current limit · 5 Capacity limit · 6 External control signal
MSG | Logs the same kind of numbered message as ERR but does not interrupt the program.
SAVE | Persists the battery's counters (per the current registration format) into a named test section; a battery can have multiple saved sections.
REST | Restores counters previously stored via SAVE (matching section name required) to continue a test from saved state; requires the matching registration format to be declared beforehand; waits or interrupts if the saved data isn't available.
TABLE — authoring | Nominal-value tables are authored in a plain-text editor (e.g. Notepad) using the pattern duration;current;power;voltage, saved as a .TXT file in the Table folder. | Pattern: X;Y;Z;U;
TABLE — loading | A table is loaded into a program step by filename. | Filename limited to 8 characters
TABLE — scaling | Factor_I / Factor_P / Factor_U scale a loaded table's current/power/voltage values for reuse across different capacities/ratings.
TABLE — built-in drive cycles | Predefined industry drive-cycle profiles usable directly, typically paired with TASK OUW for power calculation. | FUDS (fudstbl), SFUDS (sfudstabl), DST (dsttbl)
TABLE — value limiting | Optional parameters clamp table-driven values to a safe range. | Ap, An, Vp, Vn, Wp, Wn
CLEAR | Deletes all SET registrations/limits and global setpoints. Does not clear or deactivate the active registration format itself.
FILTER | Configures digital filters per channel; available on Digatron-configured hardware only.
FILE | Names the test section used for registration storage during the run. Can be used standalone, or per-cycle-iteration inside a cycle (auto-named sequentially). | Auto-naming pattern inside cycles: TSZ00001, TSZ00002, …
RANGE / IRANGE | Controls automatic vs. manual current-range switching, affecting slew rate / transit time. Can be placed anywhere in the program; takes effect at the start of the next step. | Manual values 1–4 (1–2 for IGBT); 0 = return to automatic
PARALLEL | Links multiple circuits into a synchronized parallel test; requires all linked circuits to be idle (STOP) at start. Documents specific error conditions and resolutions. | Up to 7 slaves on MBT master, up to 15 on ME master. Errors: current below set value, current exceeds max, circuit outside cabinet, circuit already in use
EIS | Runs an EIS-Meter impedance-spectroscopy step with a rich parameter set for controlling load, voltage limits, sweep frequency, and timing. | ADC (load current, sign=charge/discharge/pause) · VDC (voltage limit, requires ADC) · mHz/Hz/kHz (frequency or sweep, 8 pts/decade) · AAcMax (default 2A) · VAcMax (default 20V) · VAcMin (default -2V) · mVideal (default 10mV) · EIStime (default 10s, min 3 periods) · EISperi (default 3) · AACstart (default 20mA) · limit "n min" sets max measurement time
SYNCLine / SYNCProgram | Synchronizes circuits sharing the same SyncGroup value at a given step. SYNCProgram requires all circuits to run the identical program and version; SYNCLine does not check program identity. Inserted as a standalone, parameter-less step.

## 9. Cycles, Procedures, Counters
Cycles, Procedures & Counter/GOTO Management
Helpful program elements (section 12.5)
Feature | Description | Sub-options / Values
Cycles (BEG / CYC) | BEG marks the start of a repeatable block (optional cycle name); CYC marks its end with the repeat count as Nominal Value. Cycles can be nested/interlaced, and the repeat count can reference a SET-defined variable. | Up to 16 nested/interlaced cycles; repeat count format "N*"
Procedures — create | Combine multiple steps into a reusable named sub-program: select the lines, right-click → "Create procedure", then name it and assign a program number. Stored under the "Procedures" section of the BTS-600 program list.
Procedures — invoke | Invoked by entering the procedure's name directly in the Operator column of a program step.
Procedures — edit | A procedure is edited like a normal program. A "Show main program" toggle lets you view the calling program while a procedure step is executing.
Procedures — Decompose / Recompose | "Decompose" expands a procedure call into its inline steps for direct editing; "Recompose" collapses inline steps back into a procedure reference. Both accessed via right-click context menu.
Procedures — shared-edit warning | Editing a procedure affects every program that references it — the manual explicitly flags this as a shared-impact risk.
Counter settings (re-preset) | Cycle, Ah, and Wh counters can be re-preset via SET before resuming an aborted test, to continue tracking correctly from where the test left off.
GOTO — as Operator | Used in the Operator column with the jump destination given in Nominal Value.
GOTO — as Action | Used in the Action column with the jump destination given directly in that column, targeting a Label-defined step.
GOTO — scope restriction | Jump targets are restricted to the same program level — a GOTO cannot jump into or out of a procedure. Supports conditional branching to multiple labeled destinations.
Reserved/forbidden label names | A fixed set of names is reserved for system counters/channels and cannot be used as user-defined GOTO labels. | V, A, mA, Ah, mAh, AhCha, AhLad, mAhCha, mAhLad, AhDch, AhEla, mAhDch, mAhEla, AhStep, AhPas, mAhStep, mAhPas, AhPrev, AhPrec, mAhPrev, mAhPrec, AhStat, mAhStat, AhBal, mAhBal, Wh, mWh, WhCha, WhLad, mWhCha, mWhLad, WhDch, WhEla, mWhDch, mWhEla, WhStep, WhPas, mWhStep, mWhPas, WhPrev, WhPrec, mWhPrev, mWhPrec
Special Label — ONERROR | Jump target automatically executed when any program error occurs. Manual recommends the target sequence include a CLEAR step immediately after the jump.
Special Label — ONEXIT | Jump target automatically executed after normal program completion — e.g. used to send a completion message to the host system.

## 10. Battery Parameters
Battery Parameters
Battery-derived values usable as channel units (section 12.3)
Feature | Description | Sub-options / Values
CNom | Nominal capacity of the battery.
NoCell | Number of cells in the battery.
UGas | Gassing voltage threshold.
UMax | Maximum allowed voltage.
INom | Nominal current.
ICrank | Cold cranking current.
ChargeF | Charge factor used in balance/statistic Ah calculations.
Rin | Internal resistance.
CutOff | Final / switch-off voltage.
UNom | Nominal voltage.
EDensity | Energy density.
Usage as global limits | These parameters can be referenced as channel units directly in the Nominal Value, Limit, and Registration columns — e.g. interrupting the program whenever current exceeds INom or voltage exceeds UGas throughout the entire run.