## Slide 1
ME PROJECT · PERIPHERAL REQUIREMENTS
Primary Board — Required Peripherals
What the Primary board must provide. Digital input and output counts were read from the existing firmware, not estimated.
1
PORT
ETHERNET 10/100
Communicate with the web application. TCP/IP carries the test sequence down; UDP carries live and session data up.
Needs an external PHY and isolated connector
1
PORT
CAN FD — TEST BUS
Communicate with up to eight Secondary boards. Sends target current and voltage; receives grouped measurements.
CAN FD · 1 Mbit/s arbitration · 2 Mbit/s data · 64-byte payload
1
PORT
CAN — BMS
Communicate with the BMS network, interpreted through the CAN database uploaded from the web application.
Separate physical port from the test bus
1
PORT
CAN — EXTERNAL SENSORS
Read data from external sensors mounted around the test bench.
Separate physical port from the test bus and the BMS bus
2
PORTS
RS-485
Modbus communication. Two independent ports.
Each needs its own transceiver, biasing and termination
4
INPUTS
DIGITAL INPUTS
Interlocks and status contacts. All four are active low, and their states are reported to the web application.
PC2 / J9 · PC3 / J10 · PC6 / J8 · PC7 / J11
3
OUTPUTS
DIGITAL OUTPUTS
Contactor and indicator control. Each is driven from one bit of a control byte the web application can set.
PC17 / J10 · PC31 / J11 · PC18 / J12
1
UNIT
REAL-TIME CLOCK
Keep correct time across a power loss, without depending on the PC. Records are timestamped to 1 ms.
Requires a battery or supercapacitor backup
1
DEVICE
NON-VOLATILE MEMORY
Store EIGHT test sequences, battery and channel configuration, calibration, and power-fail recovery data.
Capacity set by the maximum size allowed per sequence
1
UNIT
WATCHDOG
Reset the board automatically if the firmware stops servicing it.
PRIMARY BOARD
the test engine
Stores and executes eight test sequences, monitors every cutoff condition, and drives eight power stages over one bus.
REQUIRED PORT AND CHANNEL COUNT
Ethernet ports
1
CAN ports
3
RS-485 ports
2
Digital inputs
4
Digital outputs
3
Real-time clock
1
Non-volatile memory
1
BTS-ME-PER-001 v3.0 · ME Peripheral Requirements · PRIMARY BOARD · Firmware Engineering · 31 July 2026
1

## Slide 2
ME PROJECT · PERIPHERAL REQUIREMENTS
Secondary Board — Required Peripherals
What each Secondary power stage must provide. Analog and digital counts were read from the existing firmware, not estimated.
1
PORT
CAN FD
The only link to the Primary. Receives targets, returns grouped readings.
CAN FD · unique node address 1 to 8 · needs termination
4
INPUTS
DIGITAL INPUTS
Interlocks and status contacts.
All four interrupt-capable
8
OUTPUTS
DIGITAL OUTPUTS
Contactor, fan and indicator control.
All eight default to OFF at reset and at power-up
1
OUTPUT
CURRENT-DIRECTION SELECT
Select the charge or discharge current path.
Defaults to a safe state at reset
1
UNIT
WATCHDOG
Reset the board if the firmware stops servicing it.
Active in every operating mode, including calibration
4
CHANNELS
HIGH-RESOLUTION ADC
1 · Battery voltage 2 · Reverse-polarity voltage
3 · Battery current, main range
4 · Battery current, second scaled range
24-bit · up to 30 kSPS · dedicated SPI
4
CHANNELS
GENERAL-PURPOSE ADC
1 · ZnT signal voltage (PC4) 2 · Battery temperature, PT100 (PA0)
3 · LnT signal voltage (PA6)
4 · Heatsink temperature
16-bit · read continuously by DMA · sensor-presence detection
1
UNIT
DAC
Sets the transistor bank drive level. The 1 ms control loop writes to it on every cycle.
20-bit · dedicated SPI · zero code means zero drive
1
DEVICE
CONFIGURATION STORAGE
Hold calibration constants and board settings.
Small — the test sequence is held on the Primary
SECONDARY BOARD
one per battery · eight on the bus
Receives a target from the Primary, regulates current and voltage to it every millisecond, and reports what it measures.
REQUIRED PORT AND CHANNEL COUNT
CAN FD ports
1
Digital inputs
4
Digital outputs
8
Analog inputs, total
8
— high resolution
4
— general purpose
4
Analog output (DAC)
1
BTS-ME-PER-001 v3.0 · ME Peripheral Requirements · SECONDARY BOARD · Firmware Engineering · 31 July 2026
2